Code No: RR-212202

Set No.

1

II B.Tech. I-Semester Regular Examinations, November-2004

ELECTRICAL AND ELECTRONICS MEASUREMENTS

(Instrumentation and Control Engineering)

Time: 3 hours							    Max Marks: 80

Answer any FIVE Questions

All Questions carry equal marks

- - -

1.	Discuss in detail the working of series type voltmeter with relevant diagram and derive the necessary expressions.

2.a)	With neat diagram explain the principle and working of AC voltmeter.

b)	An AC voltmeter calibrated for sine wave is used to measure a ramp voltage waveform rising to a peak value of 60 V in 3 m.sec. Determine the percentage error.

3.a)	Draw the typical metering circuit of a solid state VOM describing various elements used.

b)	What is the lowest full-scale voltage that could be displaced with a 100 A meter moment with an internal resistance of 150 . What would be the sensitivity of this meter in ohms /volt?

4.a)	Explain Q meter measurement of a low impedance component in the series connection.

- Explain residual inductance of a Q meter.

5.a)	Explain the relationship between trigger pulse and the sweep in an oscilloscope.

b)	Explain the block diagram of a dual time base for delayed trigger.

6.a)	Name different types of oscilloscopes. Compare their merits and demerits along with their applications.

b)	How is vertical axis of an oscilloscope is deflected? How does this differ from the horizontal axis?

7.	Describe the circuits and working principle of wave analyzers used for Audio Frequency and also in Megahertz ranges. Mention their applications and give the specifications of the instruments.

8.a)	Explain a method to measure the ratio of two frequencies using suitable block diagram.

b)	To what accuracy can a frequency counter determine an unknown frequency of       500 KHzs using a 1.2 sec time base and a time base accuracy of 0.02%.

-*-*-*-

Code No: RR-212202

Set No.

2

II B.Tech. I-Semester Regular Examinations, November-2004

ELECTRICAL AND ELECTRONICS MEASUREMENTS

(Instrumentation and Control Engineering)

Time: 3 hours							    Max Marks: 80

Answer any FIVE Questions

All Questions carry equal marks

- - -

1.	 Explain the voltmeter sensitivity and measuring method with loading effect.

2.a)	Draw the block diagram of a true RMS measuring instrument and explain its 	working.

b)	A symmetrical square wave voltage of maximum value Em and time period T is applied to an average responding AC voltmeter with a scale calibrated in terms of 	the RMS value of a sine wave. Calculate the error in meter indication.

3.a)	Explain in detail the considerations in choosing an analog voltmeter.

- A coil with a resistance of 3 is connected to the terminals of the Q meter. Resonance occurs at an oscillator frequency of 5 MHz and resonating capacity of 100 PF. Calculate the percentage of error introduced by the insertion resistance Rsh = 0.1.

4.	Explain in detail a vector impedance meter with a block diagram. Give its specifications.

5.a)	Explain the working operation of triggered time base for an oscilloscope.

b)	Why horizontal and vertical deflections are required for an oscilloscope.

6.a)	What is oscilloscope probe compensation? How is this adjusted?

b)	How does the digital storage oscilloscope differ from the conventional storage oscilloscope using a storage cathode ray tube?

7.a) 	Explain the different types of distortions caused by amplifiers.

b) 	Describe the Engineering applications of wave analyzers.

8.a)	Explain the principle of period measurement	.

b)	Explain a method of multiple period averaging scheme using suitable block diagram.

-*-*-*-

Code No: RR-212202

Set No.

3

Set No.

3

II B.Tech. I-Semester Regular Examinations, November-2004

ELECTRICAL AND ELECTRONICS MEASUREMENTS

(Instrumentation and Control Engineering)

Time: 3 hours							    Max Marks: 80

Answer any FIVE Questions

All Questions carry equal marks

- - -

1.	Derive the expression for Rs the Source Resistance for both DC ammeter and DC voltmeter and also discuss their types in detail.

2.a)	How the elements of a single phase watthour meter are connected? Explain with 	the help of a sketch.

b)	Describe the principle of measurement and working of a domestic watthour meter.

3.a)	Explain a resistance range selector circuit of a VOM.

- To check the distributed capacitance of a coil, the coil is resonated at 15 MHz with   40 PF. What is the inductance of the coil and what is the equivalent distributed capacitance.

4.a)	Name the measurements for which a vector voltmeter is used.

b)	Draw the block diagram of the vector volt meter and mention all the blocks.

5.a)	What are the advantages of dual trace over dual beam for multiple trace oscilloscopes?

b)	How is the vertical axis of an oscilloscope deflected? How does this differ from horizontal axis?

6.a)	How does the sampling oscilloscope increase the apparent frequency response of an oscilloscope?

b)	What precautions must be taken when using a sampling oscilloscope?

7.a) 	Explain the term “total harmonic distortion”. Describe the functioning of a total harmonic distortion meter.

b) 	Explain the following terms:

(i) SWR	(ii) VSWR	(iii) ISWR.

8.a)	Explain the basic principle and working of an electronic frequency counter. Also   explain how period can be measured.

b)	List the detailed specifications of electronic frequency counters.

-*-*-*-

Code No: RR-212202

Set No.

4

Set No.

4

II B.Tech. I-Semester Regular Examinations, November-2004

ELECTRICAL AND ELECTRONICS MEASUREMENTS

(Instrumentation and Control Engineering)

Time: 3 hours							    Max Marks: 80

Answer any FIVE Questions

All Questions carry equal marks

- - -

1.	Explain in detail the Permanent magnet moving coil mechanism with construction and temperature compensation.

2.a)	Give the specifications and typical values of a commercial multimeter.

b)	Design a range switch to be used in the case of an electronic voltmeter, calibrated 	in the customary range of 1-3-10 sequence with a total resistance of 12 M, to measure from 1 V to 1000 V, the meter requires 1 V for full scale deflection, at the 	gate of the FET input device.

3.a)	Give the schematic diagram of a balanced bridge DC amplifier used in electronic analog voltmeters and explain its working.

- What would true RMS reading meter indicate if a pulse wave from of 5 volts peak and a 25% duty cycle applied? What would the meter indicate if a 5volt DC input were applied (assume the meter has DC capability).

4.	Explain the working of a vector voltmeter in detail. Draw the schematic. Mention its applications.

5.a)	Explain the block diagram of a dual trace oscilloscope.

b)	Why is a delay line used in the vertical section of the oscilloscope?

6.a)	Why is an attenuator probe used? What are the advantages of using an active voltage probe?

b)	What are the advantages of digital storage oscilloscope and conventional storage oscilloscope with respect to one another?

7.a) 	What is intermodulation distortion? Describe the instrument used for the measurement of intermodulation distortion meter with the help of block diagram.

b) 	Describe the working of a ‘difference-frequency’ distortion Analyzer with the help of block diagram.

8.a)	Draw the block diagram of Versatile multiple-period measuring system and explain with suitable waveforms.

b)	List the applications of frequency synthesizers.

-*-*-*-